## Slide 1: AR-coated rCBGs 900-920nm
MIT TESTING
Arianna Bresci
02/04/2026

## Slide 2
CREOL spec sheet

## Slide 3
S45-001

## Slide 4
1mm input beam waist
rCBGs testbed
Spectral flatness looks ok, slightly transmission at 910 nm
might be due to alignment
3.01 mm
24.89 mm
12.29 mm
S45-001

## Slide 5
1mm input beam waist
FWHM 3.2 nm
FWHM 2.7 nm
FWHM 2.7 nm
S45-001

## Slide 6
5mm input beam waist
rCBGs testbed
Spectral flatness looks ok
3.01 mm
24.89 mm
12.29 mm
S45-001

## Slide 7
S45-001
5mm input beam waist
FWHM 3.7 nm
(instead of 3.2 with
1mm input beam waist)
FWHM 3.7 nm
(instead of 2.7 with
1mm input beam waist)
FWHM 3.5 nm
(instead of 2.7 with
1mm input beam waist)

## Slide 8
DC22E-080

## Slide 9
DC22E-080
rCBGs testbed
1mm input beam waist
Spectral flatness is lost below 905nm
12.21 mm
24.89 mm
6.03 mm

## Slide 10
DC22E-080
1mm input beam waist
Fit fails, too noisy output at 900 nm
FWHM 2.7 nm
FWHM 2.9 nm

## Slide 11
DC22E-080
rCBGs testbed
Spectral flatness is lost below 905nm
12.21 mm
24.89 mm
6.03 mm
5mm input beam waist

## Slide 12
DC22E-080
5mm input beam waist
FWHM 4.6 nm
(instead of 2.7 with
1mm input beam waist)
FWHM 4 nm
(instead of 2.9 with
1mm input beam waist)
FWHM 4 nm
(instead of close to no output
with 1mm input beam waist)

## Slide 13
DC22-054

## Slide 14
DC22-054
12.21 mm
24.89 mm
14.75 mm
5mm input
beam waist
1mm input
beam waist
FWHM 5 nm

## Slide 15
S50-010

## Slide 16
S50-010
10.22 mm
12.16 mm
5.99 mm
5mm input
beam waist
1mm input
beam waist
FWHM 4.1 nm
FWHM 2.7 nm

## Slide 17
DC22-079

## Slide 18
DC22-079
10.08 mm
12.72 mm
5.26 mm
5mm input
beam waist
1mm input
beam waist
FWHM 2.4 nm
FWHM 4.7 nm

## Slide 19
DC22-079
S50-010

## Slide 20
DC22-079
S50-010
With no daisy chain &
alignment constraints:
FWHM 4.4 nm
FWHM 3.7 nm
The alignment should be adjusted to a different axis to get the daisy chained VBGs to output the desired wavelengths. Experimentally more complex and no finely tuned rotation control.
The S50-010 VBG optimized for a 910nm output displays a 923.7nm output when it aligns with the DC22-079 VBG tuned to output its optimal wavelength i.e. 900nm.
5mm input beam waist